COMMONWEALTH of PENNSYLVANIA

DEPARTMENT of ENVIRONMENTAL PROTECTION

Division of Air Quality Monitoring
12th Floor
Rachel Carson State Office Building
P.O. Box 8468
Harrisburg, PA 17105-8468

Specifications for Federal Equivalent Method Ozone Analyzer
NO SUBSTITUTION PROCUREMENT

| Model# | Specifications |
| --- | --- |
| Teledyne API T400 UV Absorption Ozone Analyzer | Product and configuration must meet U.S. EPA Automated Equivalent Method: EQOA-0992-087 User Defined Configurations Range 0-100 ppb / 0-10 ppm (DUAL RANGE) Output A1: 0-100 ppb = 0-5VDC, Output A2: 0-10 ppm = 0-5VDC AC input voltage 110 Volts at 60 Hz Internal Pump Calibration Options Ambient Zero and Ambient Span Other Specifications US EPA certified operating range of 5-40°C Technical Lower detection limit <0.6 ppb Zero noise <0.3 ppb root mean square (RMS) Span noise <0.5% of reading (RMS) above 100ppb Precision <0.5% of reading above 100ppb Linearity 1% of full scale Zero drift for 24 hours <1ppb. Span drift for 24 hours <1% of reading Rise and Fall time (to 95%): <20 seconds Sample flow rate <=0.8 liters/minute One-year manufacturer warranty Internal programmable zero, span and precision points Front accessible 47mm diameter particulate filter Product manuals on CD 7", XVGA, Color LCD Graphical User Interface with touch screen I/O includes 8 isolated digital outputs, 6 isolated digital inputs Ethernet Port, 2 RS-232 ports, USB device ports Front Panel USB ports for peripherals and firmware upgrades Shall support unique TCP/IP address Temperature and Pressure compensation Internal data logging with programmable averaging periods Adaptive signal filtering to optimize response time |

Specifications for CAPS True NO2/NO/NOx Sensor
NO SUBSTITUTION PROCUREMENT

| Model# | Specifications |
| --- | --- |
| Teledyne API N500 True CAPS NO2/NO/NOx Sensor | User Defined Configurations Large, vivid, and durable color touchscreen display Power: 110W; 100-250VAC (50-60Hz) Operating Range: 0 - 40°C Calibration Options Ambient Zero and Ambient Span Other Specifications Range Min: 0 - 5 ppb full scale, Range Max: 0 - 1,000 ppb full scale Technical Zero Noise: < 0.05 ppb (RMS) Span Noise: < 0.2% of reading + 50 ppt (RMS) Lower Detectable Limit: < 0.1 ppb Zero Drift: < 0.2 ppb / 24 hours Span Drift: <0.5% of reading/24 hours Response Time: < 60 seconds to 95% Linearity: 1% of Full Scale Precision: 0.5% of reading above 5 ppb Sample flow rate: 1000 cc/min ±10% Two-year manufacturer warranty Lifetime Technical Support NumaView™ Software Long-life Particulate Sample Filter Product manuals on CD I/O includes 1 Ethernet (TCP/IP) 1 RS232 or RS485 (user-selectable) 2 Front panel USB device ports Internal data logging with programmable averaging periods |

Specifications for Dynamic Dilution Calibrator
NO SUBSTITUTION PROCUREMENT

| Model# | Specifications |
| --- | --- |
| Teledyne API T700U Dynamic Dilution Calibrator | User Defined Configurations Precise calibration for O3, NO, NO2, CO, Hydrocarbons, H2S, and SO2 in addition to other, user defined gases AC input voltage 110 Volts at 60 Hz Output voltage 0-5VDC Ozone generator with optical feedback and GPT mixing chamber UV photometer module Dilution MFC 0-20 LPM Source Gas MFC 0-100 cc/min Source Gas MFC 0-20 cc/min Other Specifications Technical – Dilution system Flow Measurement Accuracy: ± 1% of full scale Repeatability of Flow Control: ± 0.2% full scale Linearity of Flow Measurements: ± 0.5% full scale Flow Range of Dilution Air: 0 to 20 SLPM Flow Range of Cylinder Gases: 0 to 100 cc/min and 0 to 20 cc/min Zero Air Requirements: 20 SLPM @ 30 psi Input Pressure: 20-40 psig Calibration Gas Input Ports: 4 (configurable) Diluent Gas Input Ports: 1 Technical – O3 generator Maximum Output: 6 ppm LPM Minimum Output: 20 ppb LPM Response Time: 180 seconds to 98% Optical Feedback: Standard Technical – UV Photometer Range: 0-100 ppb to 0-10 ppm (selectable) Precision: 1.0 ppb Linearity: 1.0% of reading Rise/Fall Time: <20 seconds (photometer response) Response Time: 180 seconds to 95% (system response) Zero Drift: < 1.0 ppb/24 hours Two-year manufacturer warranty Product manuals on CD 7", XVGA, Color LCD Graphical User Interface with touch screen 12 digital control outputs, 12 digital control inputs 8 digital status outputs RS-232 ports Ethernet port Front Panel USB ports for peripheral devices and firmware upgrades 12 independent timers for sequences, nested sequences up to 5 levels Software linearization of mass flow controllers |

Specifications for Federal Equivalent Method Hydrogen Sulfide Analyzer
NO SUBSTITUTION PROCUREMENT

| Model# | Specifications |
| --- | --- |
| Teledyne API T101 Ultraviolet Fluorescence Hydrogen Sulfide Analyzer | User Defined Configurations Ranges: 0-50 ppb to 0-10 ppm H2S / 0-20 ppm SO2, user selectable DC Output Voltage: 0-5VDC AC input voltage 110 Volts at 60 Hz Internal pump Calibration Options Ambient Zero and Ambient Span Other Specifications Operating range of 5-40°C Technical Lower Detection Limit < 0.4 ppb Zero Noise <0.2 ppb root mean square (RMS) Span Noise <0.5% of reading (RMS) above 50 ppb Precision 0.5% of reading > 50 ppb Linearity 1% of full scale Zero Drift for 24 hours < 0.5 ppb. Span Drift for 24 hours < 0.5% of full scale Rise and Fall Time (to 95%) <120 seconds Sample flow Rate 0.65 liters/minute +/- 10% Two year manufacturer warranty and free technical support for life of analyzer Internal programmable zero, span and precision points Front accessible 47mm diameter particulate filter Product manuals on CD 7", XVGA, Color LCD Graphical User Interface with touch screen I/O includes 8 isolated digital outputs, 6 isolated digital inputs Ethernet Port, 2 RS-232 ports, USB device ports Front Panel USB ports for peripherals and firmware upgrades Shall support unique TCP/IP address Temperature and Pressure compensation Internal data logging with programmable averaging periods Adaptive signal filtering to optimize response time |

Specifications for High Performance Zero Air System
NO SUBSTITUTION PROCUREMENT

| Model# | Specifications |
| --- | --- |
| Teledyne API T701H High Performance Zero Air System | User Defined Configurations High Performance scrubbers for SO2, NO, NO2, O3, H2S, CO and Hydrocarbons AC input voltage 110 Volts at 60 Hz 30 SLPM Pump Dewpoint Sensor/Alarm Other Specifications Technical Output flow rate: 30 SLPM at 35 psig Max delivery pressure: 35 psig Dewpoint: -30C up to 30 SLPM / -45C up to 5 SLPM Dryer: Regenerative heatless dryer with lifetime > 5 years Output Concentration SO2 <0.025 ppb H2S <0.025 ppb NO <0.025 ppb NO2 <0.025 ppb O3 <0.3 ppb CO<10 ppb Hydrocarbons <0.25 ppb Compressor: Internal long life, oil-less piston pump Bench-type mounting Automatic Water Drain Modbus Connectivity Dewpoint sensor Manufacturer two-year warranty Product manuals on CD |

1. Instructions to Bidders

a. The Department of Environmental Protection (DEP), Bureau of Office Services, is requesting price bids for the following items listed in the Invitation for Bid (IFB) item list section. The bidder must supply a price for all items. Award will be made based on the lowest grand total price.

2. Delivery

a. Physical delivery location: DEP, Bureau of Laboratories, 2575 Interstate Drive, 2nd Floor, Harrisburg, PA 17110-9332, ATTN: Don Torsello.

b. Please include all costs including shipping/handling within the costs of the line items supplied, i.e. bid prices must be FOB destination. No additional costs will be allowed.

3. Agency Contacts

a. Any technical questions should be directed to Don Torsello, phone: 717-346-8354 email: dtorsello@pa.gov. Email communication is preferred.

b. Questions regarding bidding procedures should be directed to Terrell Chisholm,

phone: 717-772-5693, email: techisholm@pa.gov. Email communication is preferred.

# 4. Bid Submission

  - All Bids must be submitted electronically through the Department of General Services’ (DGS) eMarketplace website. The eMarketplace website is located at http://www.emarketplace.state.pa.us. DEP is not responsible for the maintenance of the eMarketplace website.

  - DGS’s Supplier Service Center (https://www.dgs.pa.gov/Materials-Services-Procurement/Supplier-Service-Center/ Pages/default.aspx) is available to assist vendors with registration, bidding, and account management. For questions regarding registration help, send an email to RA-PSC_Supplier_Requests@pa.gov or call (877) 435-7363, choose option 1. For questions regarding bidding help, send an email to srmhelp@pa.gov or call (877) 435-7363, choose option 2.

  - DEP is not responsible for the support or functionality of any Department of General Services website.

5. Estimated Quantities

a. The quantities contained in the IFB section are estimated only and may increase or decrease dependent upon the needs of DEP. Ancillary items may be added to this acquisition (purchase order) at the discretion of both parties at a mutually agreed upon price.

b. Payment to the contractor shall be made within approximately 30 business days after invoice receipt at the unit bid price for the quantities supplied.

6. Requirements

a. See (page 1-2 of this specification) item list section for listing of parts.

b. Teledyne is the requested manufacturer for all items contained in this bid. Substitution responses will not be accepted.

7. Payment Terms

- Payment will be rendered upon receipt of material.

8. Award

a. Award will be made to the responsible responsive bidder based on the lowest grand total price.

b. Bidders are required to complete and return the “Reciprocal Limitations Form” electronically attached to the IFB file. Reference see Title 62 Pa C.S. Section 107.