COMMONWEALTH of PENNSYLVANIA

DEPARTMENT of ENVIRONMENTAL PROTECTION

Office of Waste, Air, Radiation and Remediation 
Rachel Carson State Office Building
P.O. 2063
Harrisburg, PA 17105-8468

SILICON-CELL PYRANOMETER

NO SUBSTITUTION PROCUREMENT

| Model | Specifications |
| --- | --- |
| AIO2 Radiometer w/mount | Power Supply - 5 to 24 VDC with a nominal current draw of 300 uA Output (Sensitivity) - 1.25 mV per W m-2 Calibration Factor (Reciprocal of Output) - 0.8 W m-2 per mV (reciprocal of sensitivity) Calibration Uncertainty - < 3% at 1000 W m-2 Measurement Repeatability - < 1% Long-Term Drift - < 2% per year Non-Linearity - < 1% up to 2000 W m-2, maximum radiation measurement is 2000 W m-2) Response Time - < 1 ms Field of View - 180° Spectral Range - 360 to 1120 nm Directional Response - +/- 5% at 75° zenith angle Temperature Response- 0.04 +/- 0.04% per C Operating Environment- -40° to 70° C; 0 to 100% Relative Humidity Dimensions - 30.5 mm diameter, 37 mm height Mass - 140 g Cable- 10 m shielded, twisted-pair, santoprene rubber jacket |

1. Instructions to Bidders

a. The Department of Environmental Protection (DEP), Bureau of Office Services, is requesting price bids for the following items listed in the Invitation for Bid (IFB) item list section. The bidder must supply a price for all items. Award will be made based on the lowest grand total price.

2. Delivery

a. Physical delivery location: Attention: 2575 Interstate Drive, Harrisburg, PA 17110

b. Please include all costs including shipping/handling within the costs of the line items supplied, i.e. bid prices must be FOB destination. No additional costs will be allowed.

3. Agency Contacts

a. Any technical questions should be directed to Robert Valent, phone: 717-439-6939 email: rvalent@pa.gov Email communication is preferred.

b. Questions regarding bidding procedures should be directed to Brian Franklin, phone: 717-787-2471, email: brfranklin@pa.gov. Email communication is preferred.

# 4. Bid Submission

  - All Bids must be submitted electronically through the Department of General Services’ (DGS) eMarketplace website. The eMarketplace website is located at http://www.emarketplace.state.pa.us. DEP is not responsible for the maintenance of the eMarketplace website.

  - DGS’s Supplier Service Center (https://www.dgs.pa.gov/Materials-Services-Procurement/Supplier-Service-Center/ Pages/default.aspx) is available to assist vendors with registration, bidding, and account management. For questions regarding registration help, send an email to RA-PSC_Supplier_Requests@pa.gov or call (877) 435-7363, choose option 1. For questions regarding bidding help, send an email to srmhelp@pa.gov or call (877) 435-7363, choose option 2.

  - DEP is not responsible for the support or functionality of any Department of General Services website.

5. Estimated Quantities

a. The quantities contained in the IFB section are estimated only and may increase or decrease dependent upon the needs of DEP. Ancillary items may be added to this acquisition (purchase order) at the discretion of both parties at a mutually agreed upon price.

b. Payment to the contractor shall be made within approximately 30 business days after invoice receipt at the unit bid price for the quantities supplied.

6. Requirements

a. See (page 1 of this specification) item list section for listing of parts.

b. Met One is the requested manufacturer for all items contained in this bid. Substitution responses will not be accepted.

7. Award

a. Award will be given to the responsible responsive bidder based on the lowest grand total price.

b. Bidders are required to complete and return the “Reciprocal Limitations Form” electronically attached to the IFB file. Reference see Title 62 Pa C.S. Section 107.